Document Approval

The pre-approval of this URS Addendum will be a shared responsibility of the following functional areas:

| Functional Area | Name | Sign | Date |
| --- | --- | --- | --- |
| Created by: |  |  |  |
| Production Supervisor | Clarissa Henry |  |  |
| Reviewed by: |  |  |  |
| Production Manager | Merina |  |  |
| IT | Ade Irawan |  |  |
| QA Supervisor | Vany Okky Safitri |  |  |
| Approved by: |  |  |  |
| Site Division Head | Hadi Suryo Sutanto |  |  |
| Acknowledged by: |  |  |  |
| Quality Operation Division Head | Agus Triyantoro |  |  |

DOCUMENT CHANGE HISTORY

| Revision | Description of Change | Prepared by | Date |
| --- | --- | --- | --- |
| 00 | New Document | Triadi Wiryadinata | 13 Jan 2025 |
| 01 | Addendum | Clarissa Henry |  |

# INTRODUCTION

- Background

PT. Saka Farma Laboratories is a pharmaceutical company in Indonesia that produces pharmaceutical products such as over-the-counter medicines and health supplements. In accordance with regulations, it is necessary to ensure product quality before the product is released to the market. Quality Assurance is the totality of all arrangements made with the aim of ensuring that drugs are produced with quality according to their intended use. Quality assurance includes Good Manufacturing Practices (GMP) plus other factors. GMP aims to guarantee that drugs are made consistently, meet the specified requirements and are in accordance with their intended use. GMP includes production and quality control.

Supervision during the production process (in process control) is important in ensuring product quality. Supervision during the production process (in-process control) really needs to be done to maintain the quality of the pharmaceutical preparations that are made.

Currently, the IPC process for testing packaging integrity (Leak Test) requires digitization. Following the successful implementation of the E-IPC system in Line 1, the system will be duplicated and expanded to accommodate the production flow in Line 2 and Line 3. The process of monitoring and printing the test results needs an integrated approach. Test results are traditionally recorded limited to batch records; to see trends in test results it is necessary to process data again and process data analysis which has the potential not to be carried out in real time. Therefore, it is necessary to utilize digital application or web-based tools to check, analyze and print test results in real time across multiple production lines.

- Objectives

- Digitizing the online system for testing the leak integrity of packaged products (strips), that can be monitored, analyzed and printed in real time, with an expanded scope covering Production Line 2 and Line 3.

- The report generated from the e-Lab IPC Leak Tester can be used as a report which is attached to a batch record.

B. SYSTEM OVERVIEW

- Scope

E-IPC Leak Tester is implemented in the packaging IPC process at the Production Department of PT Saka Farma Laboratories, expanding the scope to include Line 2 and Line 3. The system must accommodate the products manufactured in Line 2 and Line 3. The system is accessible to authorized personnel (such as Operators, Supervisors, and Production Manager) in accordance with defined user roles and access mapping. The system encompasses the entire process flow, ranging from initial data entry, to report generation and audit trails, ensuring compliance with CPOB (GMP) standards and Data Integrity requirements.

- Flow Process

In general, the process flow of the E-IPC Leak Tester system operates as follows:

Module Execution:

- 20 Strips (AW & AK): The User inputs sample data (Product Code, Process Step, Batch Number). The system records the test results for the 20 strips. The results can be printed in a table, including the result status (MS or TMS), test parameters, and time stamps.

- 20 Strips per 30 Minutes: The User inputs sample data (Product Code, Process Step, Batch Number). The system records the test results for the 20 strips. The results can be printed in a table, including the result status (MS or TMS), test parameters, and time stamps.

C. REQUIREMENTS

- Operational Requirements

  - Functional Requirements

This URS is divided into several ranking levels according to the requirements of each area in relation to its business processes, as follows:

- I : Essential (the availability of this feature in the computer system may affect the continuity/smooth running of business processes)

- D : Preferred (the availability of this feature is expected to support activities/processes)

| No | Steps | Required Information | Data Source | Requirement | GMP | Rank |
| --- | --- | --- | --- | --- | --- | --- |
| Process Steps |  |  |  |  |  |  |
| 1. | 20 Strips AW & AK | Product Code, Process Step, Batch Number | Input | - The User inputs sample data (Product Code, Process Step, Batch Number). The system records the test results for the 20 strips. The results can be printed in a table, including the result status (MS or TMS), test parameters, and time stamps | Y | I |
| 2. | 20 Strips per 30 Minutes | Product Code, Process Step, Batch Number | Input | - The User inputs sample data (Product Code, Process Step, Batch Number). The system records the test results for the 20 strips. The results can be printed in a table, including the result status (MS or TMS), test parameters, and time stamps | Y | I |

- Technical Requirements

    - Hardware specification refers to SOP-IT-G004.

    - Software maintenance is easy to perform and has a guarantee period.

    - Device connected via LAN cable/WiFi of PT. Saka Farma Laboratories and/or One Kalbe.

    - Users who access according to the approved/qualified users.

- Installation Requirements

    - Web-based application, access through website.

- Account Requirements

    - The system can be accessed by personnel who have access with a username and password for each personnel.

    - One username cannot login on more than one device in same time. Login another device with the same user name will auto logout the other devices.

    - System can be access on pc’s and or tablets.

    - Login must have a forgotten password mechanism and a password notification will automatically be sent to the relevant user's email (company email).

    - System shall prohibit deletion of user information registered to the software.

    - System shall prohibits login in system failure.

    - Password must contain at least 6 characters with 3 of those characters being uppercase, lowercase, numbers, or symbols.

    - Old password cannot be duplicate when passwords are changed.

    - User lockout shall able if user fails to login for 6 attempts.

    - If user fails to login in the set number of attempts, user can login again after 15 min passed.

    - System administration shall able to register, or change users account and other system administer.

    - System administration able to give notification if there is a connectivity issues.

- Data Requirements

    - System can received data from Leak Tester in PT. Saka Farma Laboratories.

    - System shall able to view previous historical data.

    - System provide real time data and also backup database for track record based on SOP CIT.

    - System shall able to retain backup data for 5 years or more and with a minimum storage memory of 200 gb or more.

- Report Requirements

    - Download data by date within a certain period in the format pdf as per parameter: 
1. Batch number
2. Product code
Download data for analysis purposes can be in excel format, but only Supervisor can download with this format.

    - Report must consist additional e-sign from operator (tested by) and supervisor (released by).

    - Report must contain document-controlled number related to SOPs.

    - When downloaded, the report is automatically renamed based on its contents (For example: IPC MA_STFTGF64982_Lot 2; IPC Hardness_STFTGF64982_ST2_AK).

- Security Requirements

    - System shall able to restrict copying and deletion of data files.

    - System shall able to prohibit editing of data information.

    - System shall able to prohibit setup policy and change rights.

    - System shall require entry of a reason for security policy change or when a change is made in user administration such as the addition of a user account or a change to user rights.

    - Access system must use with PT. Saka Farma Laboratories intranet connection.

    - Login must be separated by user level functional (administrator, engineering (manager, supervisor, technician) Production (manager, supervisor, operator).

    - There is some access level in operational and maintenance. Input or change data record in the database available on audit trail.

    - Authority user manager level edit, and audit trail; supervisor level input, view, and audit trail; operator level input; administrator level input, edit, and audit trail.

- Performace Requirements

    - Production record should provide any information like product name, batch number, IPC test start and stop date and time, IPC test duration and process record.

    - Software maintenance is easy to perform and the system has a guarantee period.

- Document Requirements

    - Operation and maintenance manual (Manual Book)

    - Hardware specification

    - IT Infrastructure

- Interfaces

NA

- Non-functional attributes

NA

- Software Tools

NA

D. OTHER REQUIREMENTS

- ALCOA requirements

  - Attributable

Each data and filled column must be traceable to the individual who entered, including user identification date, and time.

| Role/Access Job Level | Input | View | Edit | Audit Trail |
| --- | --- | --- | --- | --- |
| Manager |  |  | √ | √ |
| Supervisor | √ | √ |  | √ |
| Operator | √ |  |  |  |
| Administrator | √ |  | √ | √ |

  - Legible

Data/documents back up or restore must be readable and complete, consistent with the original data.

  - Contemporaneous

The system is equipped with a timestamp that cannot be modified by the user or the IT administrator.

  - Original

Reports can be edited but must be traceable to the individual who edited, including user identification, date, and time.

  - Accurate

All changes (addition, deletion, and modification of data) are recorded in the audit trail or revision history.

# E. GLOSSARIUM

| Acronyms | Definition |
| --- | --- |
| URS | User Requirement Specification |
| IPC | In Process Control |